## Slide 1: DC Metering Moving toward Center Stage
June 10, 2023
1:00 PM – 2:00 PM
Tom Lawton

## Slide 2: DC Power – a Second Chance?
We all know that Tesla won the current wars.
But that was 125 years ago.
Will Tesla and Edison move to a split decision as we progress further into the 21st century?

## Slide 3: DC Power – A Resurgence
DC never quite went away but the uses were very limited.
We are now using DC in extremely prominent and quickly expanding applications
  Superchargers for Electric Vehicles
  Battery storage for renewables and other Distributed Energy Resources
We will hear more about both of these topics in the next two sessions this afternoon

## Slide 4: Caught Unprepared
While these applications have been years in the making, the actual usage was…….”Coming Soon” for over a decade. And in the last four years has exploded. The use and impact of DC in the past four years has been more of a Tsunami than a wave.
What are some of the issues this has created?
  No viable revenue grade DC meters five years ago
  No DC ANSI standards five years ago
  No traceable measuring devices or standards or metrology labs five years ago

## Slide 5: Providing the Tools
The first thing we (our electric utility and supporting manufacturing community) had to do was create tools and the guidance to our industry to address the metering of widespread DC usage.
This starts by determining what to produce and how to approve what is being produced and then how to test the equipment that has been approved.

## Slide 6: How we are trying to respond or even stay ahead of the coming demand
We have had an ANSI group working on DC standards for over six years
We published ANSI C12.32 on DC Metering a year and a half ago and are actively working to publish a more comprehensive update
We are working on ANSI C12.33 on Revenue Grade DC Transducers
Several meter manufacturers have been working on revenue grade meters
Metrology labs can now provide traceable energy measurement standards

## Slide 7: Where has DC Metering been?
To better understand where we are heading and what we need to address how to successful meter DC applications we need to understand the origins of DC metering.
The first electric meter in the 1870’s was a DC meter.
Much of the metering in the 1880’s and early 1890’s were DC meters
Image from original 1872 patent for Samuel Gardiner’s electric meter and a rendering of the front cover

## Slide 8: Where has DC Metering been?
However once AC won the contract to light the 1893 Worlds Fair, Westinghouse, Tesla and team leveraged that to win the contract to provide the generators at Niagara Falls. And the war was over.

## Slide 9: The First Power Station at the Falls
Edward Dean Adams power station at Niagara, with ten 5,000-horsepower Tesla/Westinghouse AC generators — the culmination of Tesla's dream. (Courtesy Smithsonian Institution)

## Slide 10: Relics from the Origins of the Grid
DC Metering remained for some isolated services in certain urban areas such as New York City and San Francisco
There are still DC meters in both cities (as well as in others) for elevators and other services.
These meters are not ANSI meters and utilities had no traceable means to calibrate these meters over the years.
They remained another obscure relic of the early days of our electrical grid that we kept, worked with, worked around and mainly forgot about as we do things like two phase power or DC generators.

## Slide 11: Not Quite Dead Yet
However, these “relics” are making a come back. There is suddenly a demand for DC power. Renewables are generating DC power and we are needing to both store and invert the power to AC so we can use the power on our system.
Suddenly people are needing to meter large scale DC energy usage and wondering why we need to constantly convert back and forth between DC and AC.

## Slide 12: Emerging Markets
Today almost everything we use could, with little or no modification, be run from DC. There are markets that are purely DC markets and would benefit from direct revenue grade DC metering
Electric Vehicle fast chargers
Renewable energy generation and settlement
Server farms
Micro grids
And most likely – the home of the near future

## Slide 13: DC Fast Chargers
350 KwH chargers are becoming common place. The marketplace is clamoring for larger, faster, more powerful ones that are being tested now for implementation in the near future. Conversion efficiencies are already above 97%.

## Slide 14: DC Meters for Super chargers
These ultra fast charging stations already have localized charging peaks into the mega watts.
The EV charging market is already a massive energy exchange market with energy charging peaks into the megawatts of energy.
With this amount of energy comes a significant amount of revenue which requires the most accurate billing possible.

## Slide 15: …….So We can Present Accurate Bills
With conversion efficiencies so high the best way to accurately meter this usage will be on the DC side.
Regulators can determine if a rating factor for the conversion efficiency will be used as a multiplier for the metered accuracy or if this will only come into play if conversion efficiencies fall below a certain threshold as we do with power factor on the AC side.

## Slide 16: DC Data Centers
The two most expensive costs in a data center are energy and space.
Data center operators have been steadily moving to DC distribution to minimize the conversions between AC and DC.
Moving to all DC distribution has the potential for a 5 to 25% energy reduction by improving the conversion efficiencies and reducing the amount of heat generated.

## Slide 17: DC Data Centers
DC also has the potential to save anywhere from a quarter to a third of the space
And once again – requiring/demanding accurate DC metering.

## Slide 18: ...and then we have the option to add renewables
And once a data center moves to DC distribution, the integration of renewable energies to further reduce ongoing operating costs becomes a feasible alternative.

## Slide 19: Microgrids
Any small business or even a home can become a Microgrid at this time given that virtually every appliance and tool can be powered by DC.
Inverting the AC to DC one time and in an efficient manner at the demarcation point between the home/microgrid and the utility.
  Inverting one time instead of for each piece of equipment can represent a 20% energy savings alone.

## Slide 20: DC – Edison May Win in the End
DC micro-grids prove effective for shared cogeneration and storage

## Slide 21: DC Meters – Tomorrow’s Cash Registers
If you are going to sell it, you have to measure it.
Real needs are driving the urgency for accurate and traceable DC metering

## Slide 22: Background
DC Energy – the definition
  Power
  Energy
  Four quantities to measure:
  Voltage, current, time, simultaneity

## Slide 23: Challenges
  Same definition as AC energy
  Voltage
    For DC voltages we are limited to resistive dividers to get from system voltages to ADC levels
      Generally this is not an issue.
  Current
    DC meters may use shunts, Hall effect sensors, or more sophisticated devices such as zero flux transducers
    Both high current and low current applications present issues at high accuracy

## Slide 24: Challenges
  Time Interval
    Neither measurement nor accuracy is an issue
  Simultaneity
    Modern sampling ADC’s (Analog to Digital Converter) make true simultaneity easy
  So why is DC metering considered so difficult?

## Slide 25: Challenges
  DC may not really be DC
  There may be 10% of AC ripple on top of the voltage at frequencies in the 60Hz to 360Hz range
  There may be high frequency saw-tooth waves in the 20kHz to 250kHz range
  Loads may vary rapidly including large instantaneous steps

## Slide 26: Devil is in the Details
  Modern AC meters use a variety of signal processing techniques to enhance accuracy.
    Any DC content of the signals is removed
    Algorithms such as FFT’s (Fast Fourier Transform’s) and digital transforms make use of the AC nature of the signals to improve accuracy
    For AC, small, inexpensive CTs are available with inherent accuracies of ±0.02 percent

## Slide 27: DC Current Measurement
  SHUNTS
    High currents shunts are generally large and produce very low signal levels
    A 1mΩ shunt in a 100A meter produces 10W with only a 0.1V signal
    At 1.5 amps the signal is only 0.0015 volts
    DC offsets can easily be larger than the signal and there is no simple signal processing trick to eliminate them
    Shunts generally have fairly high temperature coefficients

## Slide 28: Current Measurement
  Hall effect sensors
    Hall effect sensors are generally in the 1% to 5% accuracy range
    Closed loop sensors may be better than 1%
    Cost is generally high compared to shunts
    Temperature dependence may be an issue
    So we have technology that may be viable for the field and may be able to provide the 0.5 to 1.0% accuracies that we might want from a DC meter. How do we do better than this for calibration and certification purposes?

## Slide 29: Current Measurement
  Zero flux transformers/flux gates
    Can be very accurate <0.1% over a wide temperature range
    Measure AC and DC
    Are VERY expensive compared to all other techniques
    Few sources, large in size

## Slide 30: Establishing Traceability
  Make and Validate
    A primary reference standard candidate is constructed using multiple current sensors with AC/DC measurement capabilities
      Precision zero flux transformer for high current 10 – 1000A
      Precision shunt 0.1 ohm, 1 ppm TC with integrated PT100 temperature sensor

## Slide 31: Establishing Traceability
  Because the system is fully calibrated in AC and DC modes it is also capable of mixed signal measurement of any kind
    AC validation resolves any issues of timing and simultaneity

## Slide 32: Establishing Traceability
  Since there is no primary DC Energy Standard to provide Traceability we had to start from basic measurements, agree a methodology and qualify a primary standard.
  Working with NRC, NIST the Bureau of Weights and Measures and other interested agencies a methodology was developed and the first 17025 lab in the country has been accredited for DC energy measurement.

## Slide 33: Basic Tools
As an industry we now have some of the basic tools in place
  A DC metering standard, ANSI C12.32 has been published. This standard will be under constant revision as DC usage continues to expand and develop – but we have a starting point.
  The ability to provide traceable calibrations for devices trying to meet this standard. We have many new players jumping into the DC world who do not know anything about metering and have never considered metering as they develop new devices.
    We saw this with the initial Electric Vehicle chargers. The manufacturers of these devices have pushed back hard against being required to comply with any sort of accuracy standard – in large part because the devices did not initially meet any revenue grade standard
    The same has been true for other new areas such as standalone DC meters and embedded meters like those in street lights and chargers.
Now we have the ability to at least provide answers and guidance to these manufacturers as they develop their equipment

## Slide 34: And the Work Goes on
Now that we have the basics and now that DC metering is recognized as a field that must be mastered by regulatory agencies - and by utilities working to comply with these regulations (as we have on the AC side of the world for 125 years), the industry is able to forge ahead.

## Slide 35: Moving Ahead
National Bureau of Weights and Standards adopted into law after nearly a decade of debate and push back the first regulations on DC fast chargers
More and more DC meters are being developed as meters and not simply as embedded devices in DC devices
Guidance is being given on how to better meter DC energy

## Slide 36: Instead of the Wild West
Instead of operating in a vacuum, equipment manufacturers now have guidance on what will be needed from them.
Comments on DC Fast Charging Relevant Current Sensors- HB44 requires type certification at 1% for point of dispensing, indicating .5% or better sensor
Of consequence is location and physical constraints in EVSE dispensing systems for installing sensors. Relevance is meter/sensor complexity and that the data output needs to be isolated from measured signal, along with dedicated power for the “meter”.

## Slide 37: And Evolving
As the market continues to evolve the metering requirements will continue to evolve.
We now have the technology to run an entire household on DC. What are the implications of that?

## Slide 38: The DC Home or Small Business
All DC in the house
  Kitchen
  HVAC
  Bathroom
  Entertainment
  Communication
    Telecommunication
    Internet for work
  Car
  Solar
  Energy Storage

## Slide 39: Clear Advantages
Safety – class 2 touch safe wiring can be used throughout new construction or existing circuits can be used in existing homes, but converted to DC
Lower distribution cost – especially for remote locations

## Slide 40: Clear Advantages
Shorter potential outages as there is a back-up system
Significantly Greater energy efficiency due to fewer and more efficient energy conversions
  Within the house potentially no energy conversion – generate and store DC and then use DC. There may be nominal losses in the storage of the energy.

## Slide 41: What is the Utilities Role?
Why not take on our traditional role?
  Generation
  Distribution
  Metering
  The technology for the coming century
Why not provide the solar or another renewable and the storage as an optional part of the service.
Build truly distributed generation
Provide the AC back up and to smooth the demand as users grow in individual buildings

## Slide 42: And How do We Meter?
For AC we have established a precedent where the utility delivers the apparent power and the customer is billed for the active power. The utility loses the difference created by the power factor.
A meter with a bidirectional inverter included or immediately after the inverter would meter the DC used or in some small amount, generated.
Regulators can determine if a factor can be added for the conversion losses as these can be determined and will be fixed, like transformer losses.

## Slide 43: Conclusion
  DC Metering is here
  Today’s technology can deliver cost effective DC metering with accuracies and environmental performance similar to AC meters.
  DC applications will evolve and the form factors for DC Metering will evolve.
    Will we go with din rail mounted?
    Embedded?
    Return to our socket based meters for ease of installation and removal?
    Some combination of the above?
    None of the above?

## Slide 44: Conclusion
  Regardless of the answer, DC Metering is here today and is here to stay. DC will share the limelight with AC and will no longer be a relic of the past.
Edison is smiling today!

## Slide 45: Questions and Discussion
ISO 9001:2015 Certified Quality Company
ISO 17025:2017 Accredited Laboratory
Tom Lawton
President
tom.lawton@tescometering.com
TESCO – The Eastern Specialty Company
Bristol, PA
215.228.0500
This presentation can also be found under Meter Conferences and Schools on the TESCO website: tescometering.com