## Slide 1: New Power Definitions ANSI C12.31
Prepared by Perry Lawton, TESCO
TESCO Metering
For North Carolina Meter School
Tuesday, June 13, 8:45 AM

## Slide 2: Agenda
What is ANSI and why do I care?
What is ANSI 12.31 all about?
  Participants
  Definitions
  Contents

## Slide 3: What Is ANSI C12.31 and Why Do I Care?
Waveforms have never been in ANSI before.
You never see nice waveforms and sine waves in the field; you always see harmonics. AC and DC always have cross-signal presence.
If you take a complex metering site (light-rail system, e.g.)
World isn’t perfect – utilities and manufacturers have now agreed that here are some waveforms that represent the host of different conditions you would experience in the real world to express the reality

## Slide 4: What is ANSI? ANSI C12?

## Slide 5: ANSI
American National Standard Institute, Inc.
  Not a government agency
  Standards do not have force of Law
  All compliance is voluntary
  ANSI generates standards through the use of a sponsoring organization called the “secretariat” who actually publishes the Standard.
    For C12.1 NEMA (National Electrical Manufacturer’s Association) is the secretariat.
    NEMA has seven sections one of which is for Utility Products and Systems. All meter manufacturer’s and manufacturer’s involved in electric metering are members of NEMA or are eligible to be members.
    Paul Orr has been the NEMA’s secretary assigned to C12 for over ten years providing continuity to the process. NEMA Manufacturing members determine which efforts will be funded and which will not be funded.

## Slide 6: ANSI (cont.)
American National Standard Institute, Inc.
  NEMA organizes committees to propose and review standards
  Standards are republished approximately every five (5) years and abandoned if not updated or reaffirmed once every ten (10) years.
  Standards codify consensus approaches in common practice
  In recent years the ANSI committee has had to deal with breaking new ground as well as several controversial issues
    This is unusual for the committee and the committee has responded by no longer being reluctant to address them, staying more on point during contentious debates, and moving through the work required in a far more efficient manner. Another example of how AMI is changing everything in our industry.

## Slide 7: ANSI C12
All ANSI Standards related to Electric Metering are in the ANSI C12 group of Standards
C12 Main Committee
  General makeup has expanded slightly over last few years
  30 voting members with representation from three groups (no one group can have more than 40%):
    12 - Manufacturers: Meter, Socket, Test Equipment, etc.
    10 - Users: Utilities
    8 - General Interest: PUC, UL, IEEE, Consultants, etc.
  Meets twice a year in conjunction with EEI/AEIC Meter conference.

## Slide 8: ANSI C12 (cont.)
C12 Main Committee
  Has final approval for all activities on any C12 family standard
  Establishes Subcommittees (SC) and Working Groups (WG) to address various standards and issues
  Sub committees and Working Groups also meet twice a year in conjunction with EEI Transmission, Distribution and Metering Conference and also hold regular or ad hoc conference calls throughout the year as members put together drafts and other technical material for consideration at the next face to face meeting.

## Slide 9: ANSI C12 (cont.)
C12 Subcommittees
  Various subcommittees have been organized to review specific standards
  This is where the work is really done
  Each operates slightly differently
  Each meets on a schedule of its own choosing
  Most meet at EEI Biannual Transmission, Distribution and Metering Meetings
  Communication WG meets more often and longer
  Various subgroups meet frequently by teleconference

## Slide 10: ANSI C12.31 for Electricity Meters— Measurement of AC Power, Energy and Power Factor

## Slide 11: ANSI C12.31 – Who’s in it

## Slide 12: ANSI C12.31 – Interpreting it
Defines AC electrical power (active, reactive and apparent), AC electrical energy (active, reactive and apparent) and power factor in terms of sampled voltage and current measurements.
The definitions are provided to facilitate uniform comparison of the power, energy and power factor measurement values reported by electricity meters in comparison to equipment used as reference standards that implement these definitions for the determination of meter accuracy in the time domain and frequency domain.
In the limit of “perfect implementation” the answers given by the time domain and frequency domain formulas should give the same answers within the claimed uncertainty limits.

## Slide 13: ANSI C12.31 – Interpreting it
Previous ANSI C12 Standards have only defined the value for VA when the waveforms for current and voltage were sinusoidal.
Since in the modern world this is seldom the case outside of the laboratory, this standard is a first step to bring definitions and testing more in line with the contemporary load environment.
The purpose of this document is to define a standard value that can be computed from a set of simultaneous samples of voltages and currents. The exact computational method to be used is not prescribed, only that the resultant value from whatever calculational method is employed must get the same result as defined by this standard.

## Slide 14: C12.31 – What’s in it
General definitions
Single phase, single cycle equations
Quadrant-based definitions
Measurements for greater than one cycle
Polyphase definitions
Annex of informational examples and diagrams

## Slide 15: General Definitions
4.2.1 - Sampled Waveform
  Reduction of the continuous-time waveform to a set of discrete, evenly spaced amplitude values
4.2.2 - Sampling Rate (R)
  The constant rate, in samples per second, at which the discrete measurements of a signal are made.
4.2.3 - Integration Period
  The integration period is defined as precisely one cycle of the fundamental of the reference voltage waveform.
4.2.4 - Test Interval
  The total time of a measurement used in the performance of a test. The test interval may include both whole and fractional cycles.
4.2.5 - Waveform Representation
  For the purpose of this standard, voltage and current waveforms are represented as either a set of data points taken at a constant sample rate, or as a Fourier series as detailed in 4.2.5.1
Notes: Sine wave – energy is area under the curve. Each meter manufacturer defined a different way of doing it. This defines acceptable methodologies for the modeling of these real-world anomalies or special waveforms.

## Slide 16: General Definitions (cont)
4.2.5.1 – Fourier Sequence
  These equations can represent any waveform over a single cycle.
  The coefficients an and bn, cn and dn and ω0 can vary cycle to cycle.
  When these coefficients (an, bn, cn and dn) are used in calculations, it shall be understood that they are calculated based on a single cycle starting at and ending at .
  Note: Generally, Fourier analysis is applied to a data stream containing many cycles that are identical. However, here we apply the technique to a single cycle.
  For the purpose of this standard, shall be the positive going zero crossing of the fundamental of the phase A voltage.
Eq. 4.2.5.1.1
Eq. 4.2.5.1.2
Notes: Note: is positive when the current LEADS the voltage and negative when the current LAGS the voltage. This choice of sign preserves the normal Cartesian right-handed coordinate system such as the convention that positive angles are measured counterclockwise from the x-axis. It also preserves the normal way in which we draw phasor diagrams as is discussed in Annex D.
Note: 𝜽_𝒏 is positive when the current LEADS the voltage and negative when the current LAGS the voltage. This choice of sign preserves the normal Cartesian right-handed coordinate system such as the convention that positive angles are measured counterclockwise from the x-axis. It also preserves the normal way in which we draw phasor diagrams as is discussed in Annex D.

## Slide 17: Single Phase, Single Cycle
ANSI C12.31 defines single phase, single cycle equations.
Acknowledges that the equations in the standard are computed cycle to cycle, while in the real world, they can the actual period may vary cycle to cycle.

## Slide 18: Quadrant-based Definitions
For some applications it may be desirable to separate power and energy quantities into bins based on the flow direction, delivered (imported) and received (exported), of active and reactive energy.
Provides consensus framework for doing so

## Slide 19: Measurement of >1 Cycle
When the measurement period is an integer number of cycles, then the result is computed either as the weighted average or sum (depending on the quantity) of the individual cycles.
For multiple cycle intervals the beginning of one cycle is coincident in time with the end of the previous cycle.
Defines integer number of cycles, and arbitrary time intervals

## Slide 20: Polyphase Definitions
  Establishes consensus definitions for Polyphase Notation, Polyphase Energy, and Polyphase Power Factor.
Example of Figure 8.2.1.2, vector apparent energy

## Slide 21: ANSI C12.31 – What’s it do?
ANSI C12.31 provides a legitimate framework and establishes consensus for key definitions required for:
  “[…]Uniform comparison of the power, energy and power factor measurement values reported by electricity meters in comparison to equipment used as reference standards that implement these definitions for the determination of meter accuracy in the time domain and frequency domain.”

## Slide 22: Questions and Discussion
Perry Lawton
Northeast Regional Sales
Perry.Lawton@tescometering.com
TESCO – The Eastern Specialty Company
Bristol, PA
215-228-0500
This presentation can also be found under Meter Conferences and Schools on the TESCO website: tescometering.com
ISO 9001:2015 Certified Quality Company
ISO 17025:2017 Accredited Laboratory

## Slide 23: Tesco hospitality suite
We would like you to join us in the TESCO Hospitality Suite for networking and more discussions about metering. The discussion will not be exclusively metering……….but we love metering and that is the most common topic.
TESCO Hospitality Suite – Brighton Tower
Monday and Tuesday 8:00 PM – 10:00 PM

## Slide 24: Definitions
ANSI C12.24 TR:2022 Definitions for Calculations of VA, Vah, VAR, and VARh for Poly-Phase Electricity Meters
t Time,
t0 Starting point in time for a single cycle calculation, in this standard the positive going zero crossing of the fundamental of the phase A voltage
T0 Period of the cycle starting at t0 ‡
Instantaneous value of the voltage at time t
Sampled instantaneous voltage (see note 2)
N The number of samples taken during a single cycle of the reference voltage waveform
V Root mean square voltage measured over a single cycle ‡
Vn Root mean square amplitude of the nth harmonic of the voltage
Instantaneous value of the current at time t
Sampled instantaneous current (see note 2)
I Root mean square current measured over a single cycle ‡ of the voltage waveform
In Root mean square amplitude of the nth harmonic of the current
Instantaneous active power
P Active power computed over a single cycle ‡ of the voltage
P1 Fundamental only component of the active power computed over a single cycle ‡ of the voltage
Active energy computed over a single cycle ‡ of the voltage
Fundamental only component of the active energy computed over a single cycle ‡ of the voltage
Reactive power computed over a single cycle ‡ of the voltage
Fundamental only component of the reactive power computed over a single cycle ‡ of the voltage
Reactive energy computed over a single cycle ‡ of the voltage
Fundamental only component of the reactive energy computed over a single cycle ‡ of the voltage
Apparent power computed over a single cycle ‡ of the voltage
Fundamental only component of the apparent power computed over a single cycle ‡ of the voltage
Apparent energy computed over a single cycle ‡ of the voltage
Fundamental only component of the apparent energy computed over a single cycle ‡ of the voltage
= Used where a definition is based on continuous variables
≈ Used when a definition is based on sampled data (time or frequency domain), implies equality in the limit of sufficiently high sampling rate
‡ Note 1: In this document single cycle means one period (T0) of the fundamental of the reference voltage waveform. In polyphase systems it means one period of the reference phase voltage (typically phase A) waveform. In the equations in this standard the integration is shown as starting at a time and ending at