## Slide 1
ANSI Standards Review
Prepared by Bill Hardy,TESCO
For North Carolina Electric Meter School
Thursday, June 29, 2017 at 8:00 a.m.

## Slide 2: The Presenter
Member of ANSI C12 since 2004
  Voting member C12 Main, C12SC1, C12SC15, C12SC16; C12SC29
    SC16 C12.20 (Vice-Chair)
    SC29 Field Testing (Chair)
    SC31 VA Measurement Standard (Chair)
    SC46 Electricity Metering 0.1. 0.2, 0.5 Accuracy Classes for Measurement of Active, Apparent and Reactive Power (Chair)
  Author: Power Measurements Handbook
  Contributing Author: Handbook of Electricity Metering

## Slide 3: ANSI
American National Standard Institute, Inc.
  Not a government agency
  Standards do not have force of Law
  All compliance is voluntary
  ANSI doesn’t actually generate any standards
  Each standard is controlled by an industry organization as the “secretariat”
    For C12 NEMA (National Electrical Manufacturer’s Association) is the secretariat
    Paul Orr is NEMA’s secretary of C12

## Slide 4: ANSI
American National Standard Institute, Inc.
  NEMA organizes committees to propose and review standards
  Standards are republished approximately every 5 years
  Standards codify consensus approaches in common practice
  Generally, they do not break new ground or deal in controversial issues.
    This is changing. Can’t avoid issues any longer.

## Slide 5: ANSI C12
C12 Main Committee
  General makeup has expanded slightly over last few years
  34 voting members with representation from three groups:
    12 - Manufacturers: Meter, Socket, Test Equipment, etc.
    13 - Users: Utilities
    9 - General Interest: PUC, UL, IEEE, Consultants, etc.
  Usually meets twice a year in conjunction with EEI/AEIC

## Slide 6: ANSI C12
C12 Main Committee
  Has final approval for all activities on any C12 family standard.
  Establishes Subcommitties (SC) and Working Groups (WG) to address various standards and issues.
  Meets twice a year in conjunction with EEI Transmission, Distribution and Metering Conference.

## Slide 7: ANSI C12
C12 Subcommittees
  Various subcommittees have been organized to review specific standards
  This is where the work is really done.
  Each operates slightly differently
  Each meets on a schedule of its own choosing
  Most meet at EEI Biannual Transmission, Distribution and Metering Meetings
  Communication WG meets more often and longer
  Various subgroups meet frequently by teleconference.

## Slide 8: ANSI C12 – Sub-Committees
Sub-Committee | Standards
C12 SC1 | C12.1, C12.4, C12.5, C12.10
C12 SC15 | C12.6, C12.7, C12.8, C12.9, C12.11
C12 SC16 | C12.20
C12 SC17 | C12.18, C12.19, C12.21, C12.22, C12.23, C12.26
C12 SC29 | C12.29 Standard for Field Testing of Metering Sites
C12 SC31 | C12.31 VA Measurement Standard
C12 SC46 | C12.46 American National Standard for Electricity Meters - 0.1, 0.2 and 0.5 Accuracy Classes for the Measurement of Active, Apparent and Reactive Power

## Slide 9: ANSI C12 SC 1
C12.1 -2014 – Code for Electricity Metering
  Published 6/30/2016
  Changes in this Revision
    Vector diagrams for balanced voltages are defined for common services.
    Various definitions have been updated.
    Bi-directional metering is added and tests updated to include bidirectional measurements.
    Significant changes were made to the temperature rise section based on TRINIWOG tests

## Slide 10: ANSI C12 SC 1
C12.1 – Code for Electricity Metering
  Changes in this Revision (continued)
    Table 1b on maximum errors for reference standards was added.
      Tightened to 0.05% and 0.02%
    Accuracy Class 0.5 was added to the standard
    A new column was added to all tables with tighter specs for AC 0.5
    Section 5 on Standards for New and In Service Performance was completely rewritten
    Appendix D was rewritten and simplified
    A new Appendix F was added

## Slide 11: ANSI C12 SC 1
C12.1 – Code for Electricity Metering
  Safety
    Work is under way to move the safety related sections of C12.1 into C12.10
    C12 is coordinating efforts with UL and its work on Subject 2735
SC 1 is in the process of being broken into two separate committees.
  SC1 will focus on C12.1.
  The new SC will focus on C12.10

## Slide 12: ANSI C12 SC 1
C12.10 – Physical Aspects of Watthour Meters – Safety Standards
  Published C12.10 – 2011
  New revision ready for ballot
    Will add safety sections from C12.1
    New safety requirements in development
    Working with UL to achieve consistency with UL 2735C
    New meter forms with auxiliary power pins have been added

## Slide 13: ANSI C12 SC 1
New meter forms with Auxiliary Power
  Seven new forms approved
Old Form | New Form | AUX Power Pin | AUX Neutral Pin
3S | 103S | 6 | 5, 5A
9S | 109S | 14 | 17
12S | 112S | 6 | 5, 5A, 10
35S | 135S | 14 | 17
36S | 136S | 14 | 17
45S | 145S | 14 | 17
66S | 166S | 14 | 17

## Slide 14: ANSI C12 SC 1
New meter forms with Auxiliary Power
Transformer Rated
Form 109
Form 136
Form 166

## Slide 15: ANSI C12 SC 1
New meter forms with Auxiliary Power
  Transformer Rated
Form 135
Form 145

## Slide 16: ANSI C12 SC 1
New meter forms with Auxiliary Power
  Self contained
Form 103
Form 112

## Slide 17: ANSI C12 SC 15
C12.6 – Marking and Arrangement of Terminals for Phase-Shifting Devices used in Metering
  Re-Published without change Sept 2012
C12.7 – Requirements for Watthour Meter Sockets
  Published – Feb 2015
Notes: 17

## Slide 18: ANSI C12 SC 15
C12.8– Test Block and Cabinets for Installation of Self contained “A” Base Watthour Meters
  Published without change Dec 2012
C12.9 – Test Switches and Plugs for Transformer Rated Meters
  Extensive Revision just completed
  Specifications for test plugs added
  Published – March 2015
Notes: 18

## Slide 19: ANSI C12 SC 15
C12.11– Instrument Transformers for Revenue Metering, 10kBIL through 350 kBIL
  Revised to parallel C57.13
  Published – July 2014
Notes: 19

## Slide 20: ANSI C12 SC16
C12.20-2015 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Published April 27, 2017
  0.1 Percent Accuracy Class Added
    New column added to all accuracy test tables
  Section 4.6 Added to provide specifications for test outputs
    Type B port specifications were made “informative” due to last second objections from a meter manufacturer

## Slide 21: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Tighter performance at reference conditions was specified
  Table 2 updated to make clear which forms and applications are Blondel compliant.
  Table 2A added to make clear which forms and applications are NOT Blondel compliant

## Slide 22: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Polyphase loading for accuracy testing is allowed and details provided
    As of Jan 1, 2020 polyphase loading will be REQUIRED
  Non-Sinusoidal Performance Testing
    Extensive performance testing using non-sinusoidal waveforms was added

## Slide 23: Harmonic Load Waveforms
ANSI C12.20 now addresses harmonic waveforms as well as sinusoidal.
Chart: 
Chart: 
Chart: 
Chart: 
Chart: 
Chart: Voltage

## Slide 24: ANSI C12 SC 17
C12.18 – Protocol Specification for ANSI Type 2 Optical Port
    Reaffirmed and published 3/16/2017
C12.19 -2015– Utility Industry End Device Data Tables
    Major additions in process.
C12.21 – Protocol Specification for Telephone Modem Communication
    Published 3/21/17

## Slide 25: ANSI C12 SC29 Field Testing of Metering Installations
Field Testing of Metering Accuracy
  Nearly ready for publication
  New approach to testing
    Perform the test in any of several approaches
    If test is within the specified window then it is acceptable

## Slide 26: ANSI C12 SC29 Field Testing of Metering Installations
Low Temperature Range (LTR) | Tmin < T < 0°C | where Tmin is the lowest operating temperature certified by the manufacturer
Normal Temperature Range (NTR) | 0°C ≤ T ≤ 50°C | Range over which performance is expected to match the “Nominal” requirements
High Temperature Range (HTR) | 50°C < T < Tmax | where Tmax is the highest operating temperature certified by the manufacturer
Table 1 – Temperature ranges
Normal Harmonic Distortion (NHD) | THD ≤ 30% | Range over which harmonic distortion on the voltage is expected to have no significant effect.
High Harmonic Distortion (HHD) | THD > 30% | Range over which harmonic distortion on the voltage may have an effect outside of normal accuracy expectations.
Table 2 Voltage Quality (Harmonic content)

## Slide 27: ANSI C12 SC29 Field Testing of Metering Installations
Current Class | Low Current | Normal Current
2 | I < 0.025 | 0.025 ≤ I ≤ 2
10 | I < 0.25 | 0.25 ≤ I ≤ 10
20 | I < 0.25 | 0.25 ≤ I ≤ 20
100 | I < 1.5 | 1.5 ≤ I ≤ 100
200 | I < 3.0 | 3.0 ≤ I ≤ 200
320 | I < 5.0 | 5.0 ≤ I ≤ 320
Table 3 – Current ranges by current class
Low Harmonic Content (LHC) | THD ≤ 100% | Range over which harmonic distortion on the current is expected to have no significant effect.
High Harmonic Content (HHC) | THD > 100% | Range over which harmonic distortion on the current may have an effect outside of normal accuracy expectations.
Table 4 – Current ranges by harmonic content

## Slide 28: ANSI C12 SC29 Field Testing of Metering Installations
Meter Accuracy Class | Minimum Test Equipment Accuracy | Preferred Test Equipment Accuracy
0.1% | 0.025% | 0.01%
0.2% | 0.05% | 0.02%
0.5% | 0.05% | 0.05%
1.0% | 0.05% | 0.05%
Table 1 – Field Test Equipment Accuracy

## Slide 29: ANSI C12 SC29 Field Testing of Metering Installations
Meter Accuracy Class | Maximum Expected Error
0.1% | 0.25%
0.2% | 0.5%
0.5% | 1.0%
1.0% | 2.0%
Table 2 – Expected Test Accuracy

## Slide 30: ANSI C12 SC29 Field Testing of Metering Installations
Issues still to be addressed
  Field Testing of Current Transformers in the Metering Circuit
    Ratio Testing
    Ratio Testing with Applied Burden
    Admittance Testing
  Field Testing of Potential Transformers in the Metering Circuit
    Ratio Testing
    Ratio Testing with Applied Burden
  Validation of Wiring

## Slide 31: ANSI C12.31
At the moment there is no non-sinusoidal definition for VA
New ANSI Standard coming very soon
C12.31

## Slide 32: New Definitions
RMS Voltage
Basic Definition
Time Domain
Frequency Domain
Waveform

## Slide 33: New Definitions
RMS Current
Basic Definition
Time Domain
Frequency Domain
Waveform

## Slide 34: New Definitions
Active Power
Basic Definition
Time Domain
Frequency Domain

## Slide 35: New Definitions
Apparent Power
Basic Definition
Time Domain
Frequency Domain

## Slide 36: ANSI C12.46
New Sub-Committee to Develop a Replacement for C12.1 and C12.20
  Adopts the structure of OIML IR-46
  Addresses 0.1, 0.2, and 0.5 Accuracy Classes
  Addresses Active, Reactive, and Apparent Energy
  Will be biggest change to metering standards in 100 years
  Review and editing of draft in process
  Expected Publication 2020

## Slide 37: Questions and Discussion
Bill Hardy, CTO
TESCO – The Eastern Specialty Company
Bristol, PA
865-279-1090 (cell)
215-785-2338 (office)
This presentation can also be found under Meter Conferences and Schools on the TESCO website: www.tesco-advent.com