## Slide 1
January 28, 2019
Atlanta, GA
Introduction to
Self-Contained
Metering
Prepared by Tom Lawton, TESCO
The Eastern Specialty Company
for PREA Metering School
March 10, 2020
8:15 a.m. – 9:30 a.m.
Notes: 1

## Slide 2: Questions to Answer
Why do we test?
How do we test?
What types of meter tests are there?
How do utility tests differ from customer request tests?
What is In-Service Testing?
How do we know meter tests are good?
What do we do with the test data?
What is the electrical theory behind the testing we do?

## Slide 3
Our regulatory commissions require us to test meters.
But only for accuracy. State regulatory commissions want electric utilities to ensure that no customer is being billed unfairly and that no subset of customers is being unfairly subsidized by the rest of the rate payers. Some states mandate only accuracy tests and others require demand and time of use accuracy tests.
Any tests beyond accuracy tests are tests that are simply good business practice.
Why Do We Test?

## Slide 4
Complaint Testing
Customers always have the right to request a meter test.
Some utilities and some jurisdictions allow for testing at the customer site, others require a test in a laboratory environment.
Some allow the customer to witness the test and others require the utility commission to witness the test.
Utilities must show that the meter tests well and must demonstrate that they have a test program in place to ensure the meters in service are performing well.

## Slide 5: General Meter Testing Requirements
New Meters
  Manufacturers tests
  In-house tests on new shipments
Return to Service Testing
In-Service Meters
  Periodic Tests
  Selective, random, or statistical testing
Retirement tests
Testing of related metering equipment

## Slide 6
New Meter Testing Programs
Accept the Manufacturer’s Test results
Perform a Statistical Test of an incoming shipment
Perform a 100% test of an incoming shipment

## Slide 7
Return to Service Testing
Meters to be returned to service must always (virtually every utility commission requires this) be accuracy tested before being returned to service.
Best business practices also require that the meter is functionally tested as well.

## Slide 8
In Service Testing
Meter Testing for new and in-service meters is specified in ANSI C12.1-2015, American National Standard for Electric Meters, Code for Electricity Metering. Most utility commissions use this Standard a reference or the basis for their meter testing requirements.

## Slide 9: Test Plans for Meters
Four test plan options available:
Periodic
Statistical
Notes: UMPIRE is a
Universal Meter Programming and Reading System
It reads all meters and recorders
It can control all manufactures software

## Slide 10: Periodic Test Plans
Periodic
  Varies by State
  Example provided by ANSI C12.1:
    Each Electro Mechanical meter is tested once every 8 years
    All other Meters are tested every 16 years
    Appendix D provides details for other meters & devices
    No guidance for AMI meters
  Generally, average of 12.5% of population tested per year
8 Years
16 Years

## Slide 11
Statistical Test Plans The Best Approach
ANSI C12.1-2001 Code for Electricity Metering Guidance
Paragraph 5.1.4.3.3 Statistical sampling plan
“The statistical sampling plan used shall conform to accepted principles of statistical sampling based on either variables or attributes methods. Meters shall be divided into homogeneous groups, such as manufacturer and manufacturer’s type. The groups may be further divided into subdivision within the manufacturer’s type by major design modifications.”
NOTE - Examples of statistical sampling plans can be found in ANSI/ASQC Z1.9, the ANSI version of MIL-STD-414 and ANSI/ASQC Z1.4, the ANSI version of MIL-STD-105.

## Slide 12
Why Use a Statistical Testing Plan?
Focuses testing on the proper meters
Minimizes number of meters to be tested; usually requires less than 30% of what a periodic testing plan requires
Provides data and analysis tools for use in
understanding what is happening with installed meters
or for use in the purchasing of new meters

## Slide 13
Homogeneous Population(s)
The groups or populations being sampled and tested
are made up of the same or similar items, items
which operate in the same way and were made in the
same manner.
For electric meters, this has traditionally been
interpreted as being meters of a specific meter
type from a manufacturer (i.e. AB1, J5S, MX, etc.).
AMR & AMI programs have helped to make the
overall populations more homogenous. This makes a
utility with AMR & AMI meters better prepared to
take advantage of a statistical sampling plan.

## Slide 14
Testing of a Meter vs Testing a Site
Test an installation and system and not just a meter!
Test programs may need to involve testing and checking the meter performance as well as checking and testing the installation. This more extensive test check list needs to be done especially for the higher revenue C&I customers.

## Slide 16: Meter Testing Traceability
Test equipment to NIST standards
Tracking number of meters to be tested per State Commission requirements
Tracking meter test data
    Meter Records
    Meter Data Management System (MDMS)

## Slide 17
Traceability is defined as ability to link the results of the calibration and measurement to related standard and/or reference (preferably national or international standard) through an unbroken chain of comparisons.
Calibration is typically performed by measuring a test unit against a known standard or reference.
Master standard (i.e. gages) are kept by National Measurement Institute (NMI) of each country.
Meter Testing Traceability
Notes: 17

## Slide 18
National Institute of Standards and Technology (NIST) provides internal tracking numbers, which are often used as evidence of traceability.
WARNING! NIST does not certify or guarantee that calibration and measurements are correct, nor does it provide any kind of certification of accuracy and calibration. NIST only provides certifications for the work performed by them.
Meter Testing Traceability
Notes: 18

## Slide 19
Meter Testing Traceability - Standards
Notes: 19

## Slide 20
National Standard
In the US, this is maintained by NIST, in Canada by NRC. Not all countries have a National Standards group and even the US does not have a group for every item of interest to an electric utility (e.g. voltage transformers)
Reference/Master Standard
Item of highest metrological quality located at a site where calibration is being conducted.
Transfer Standard
Lower level of Reference Standard and used for calibration of lower level calibration requirements measuring devices.
Working Standard
  Lower level of Reference Standard and used for calibration of lower level calibration requirements measuring devices. Should be compared to Master Standard or Reference Standard on regular basis; used for daily checks comparisons of the calibrated devices.
Meter Testing Traceability - Standards
Notes: 20

## Slide 21
Test Equipment Calibration
Primary Requirement: Traceable to NIST Standards
Meter Test Boards, Field Test Kits calibrated to a known master standard maintained at Meter Shop.
  Some periodicity such as monthly or quarterly
Reference or Master standard calibrated by outside vendor traceable to NIST or directly by NIST.
  Usually annually

## Slide 22: Tracking Meter Records
AMI programs help to update and overhaul meter record systems.
Having the records for the entire meter population updated allows for a better chance that test data is available to answer questions and that any meter may be selected as part of the sample for testing.

## Slide 23
Meter Test Data Tracking
Test data should be tracked throughout meter life
  Certification testing, first article, acceptance testing, in-service (field & shop), retirement
Meter test data should be linked to meter record data such as meter form, amps, voltage, display type, etc.
Best time to start to develop the program is before the meters are being installed.
Accuracy test data is usually collected automatically as new meters are tested in meter shops or cross docks.

## Slide 24
Meter Test Data Tracking
Need to consider tracking non-accuracy functional testing (meter software configuration, service disconnect testing, voltage, etc.)
Use installation reports to determine if there is any initial concerns about the meters being installed.
Typical reports that should be available:
  Failed Meter Report, Project to Date
  Electric Meters on Network Report

## Slide 25
Meter Test Data Tracking System
System should track meter test results for ease of future reference or for response to public or Utility Commission inquiries.
Maybe part of Meter Data Management System (MDMS) or a separate Meter Records system.
Requires discipline in collecting & entering data, especially field tests.

## Slide 26: Basic Electrical Laws
Ohms Law
  Voltage = Current times Resistance
  V = I X R
  THE MOST USEFUL AND THE MOST FUNDAMENTAL OF THE ELECTRICAL LAWS

## Slide 27: Basic Concepts: Electricity and Water
Comparing Electricity to Water flowing from a hose
  Voltage is the equivalent of the pressure in the hose
  Current is water flowing through a hose (coulombs/sec vs gal/sec). The water in a system is the “charge” (coulombs)
  Impedance(Resistance) is the size of the hose. The nozzle would provide a change in resistance.
  Power is how fast water flows from a pipe (gallons per minute vs kilowatts). Power is a rate of energy consumption

## Slide 28: Basic Concepts: Electricity and Water

## Slide 29: Practical Electricity

## Slide 30: Basic Electrical Laws
Ohms Law Examples
  If V = 20 volts and I = 5 amperes what is the resistance?
    R = V / I = 20 / 5 = 4 ohms
  If R = 20 ohms and V = 120 volts what is the current?
    I = V / R = 120 / 20 = 6 amps
  If I = 10 amperes and R = 24 ohms what is the voltage?
    V = I x R = 10 x 24 = 240 volts
  Problem: If V = 240 volts and R = 6 ohms what is the current?
    I = V / R = 240 / 6 = 40 amps

## Slide 31: Basic Electrical Laws

## Slide 32: Basic Electrical Laws
Power = Voltage x Current = V x I = I2R = V2/R
  If V = 20 volts and I = 8 amperes what is the power?
    P = V x I = 20 x 8 = 160 watts
  If R = 5 ohms and V = 120 volts what is the power?
    P = V2/R = 120 x 120 / 5 = 2880 watts
  If I = 10 amperes and R = 20 ohms what is the power?
    P = I2R = 10 x 10 x 20 = 2000 watts
    1 kilowatt (kW) = 1,000 watts
    1 megawatt (MW) = 1,000,000 watts

## Slide 33: Basic Electrical Laws
Kirchoff’s Voltage Law (KVL)
The sum of the voltages around a circuit loop is zero.
  V = V1 + V2 + V3 = I * (R1 + R2 + R3) = I1 * R1 + I2 * R2 + I3 * R3
  I = I1 = I2 = I3
Resistors in series add

## Slide 34: Basic Electrical Laws
Kirchoff’s Voltage Law (KVL) – PROBLEM #1
  I = V / R = V / (R1 + R2 + R3) = 120/(10+20+30) = 2 amperes
  V1 = I*R1 = 2 *10 = 20 volts
  V2 = I*R2 = 2 *20 = 40 volts
  V3 = I*R3 = 2 *30 = 60 volts
What is the current in the circuit?
What are V1, V2, V3?

## Slide 35: Basic Electrical Laws
Kirchoff’s Current Law (KCL)
The sum of the currents at a node in a circuit is zero.
  I = I1 + I2 + I3
  If Loads are placed in parallel they sum up.
  So does power
  P = V x I = V x I1 + V x I2 + V x I3

## Slide 36: Basic AC Theory Power – The Simple View
E = Voltage (rms)
I = Current (rms)
PF = Power Factor
Power = Watts = E x I x PF
Power is sometimes
referred to as Demand
Sinusoidal
Waveforms
Only
NO
Harmonics
For a 120 Volt service drawing
13 Amps at Unity (1.0) PF,
how much power is being drawn?
Power = 120 x 13 x 1.0 = 1560 Watts

## Slide 37: Basic Meter Math Power – The Simple View
For a 120 Volt service drawing
13 Amps at 0.866 PF,
how much power is being drawn?
Power = 120 x 13 x 0.866 = 1351 Watts
For a 480 Volt service drawing
156 Amps at 0.712 PF,
how much power is being drawn?
Power = 480 x 156 x 0.712 = 53,315 Watts

## Slide 38: Basic AC Theory Power – The Simple View
In the previous example we had:
Power = 480 x 156 x 0.712 = 53,315 Watts
Normally we don’t talk about Watts, we speak in Kilowatts
1000 Watts = 1 Kilowatt = 1 kW
Watts / 1000 = Kilowatts
For a 480 Volt service drawing
156 Amps at Unity (0.712) PF,
how many Kilowatts are being drawn?
Power = 480 x 156 x 0.712 / 1000 = 53.315 kW

## Slide 39: Basic AC Theory Energy – What We Sell
If power is how fast water flows from a pipe,
then energy is how much water we have in a bucket
after the water has been flowing for a specified time.
Energy = Power x Time
1 kW for 1 Hour = 1 Kilowatt-Hour = 1 kWh
Energy (Wh) = E x I x PF x T
where T = time in hours
Energy (kW) = (E x I x PF / 1000) x T

## Slide 40: Basic Meter Math Energy – What We Sell
For a 120 Volt service drawing 45 Amps at a Power Factor of 0.9 for 1 day,
how much Energy (kWh) has been used?
Energy = (120 x 45 x 0.9 / 1000) x 24 = 116.64 kWh
For a 240 Volt service drawing 60 Amps at a Power Factor of 1.0 for 5.5 hours,
how much Energy (kWh) has been used?
Energy = (240 x 60 x 1.0 / 1000) x 5.5 = 79.2 kWh

## Slide 41: Basic Meter Math Energy – What We Sell
For a 120 Volt service drawing 20 Amps at a Power Factor of 0.8 from 8:00AM to 6:00PM,
and 1 Amp at PF=1.0 from 6:00PM to 8:00AM
how much Energy (kWh) has been used?
8:00AM to 6:00PM = 10 hours
6:00PM to 8:00AM = 14 hours
Energy = (120 x 20 x 0.8 / 1000) x 10 = 19.2 kWh
Energy = (120 x 1 x 1 / 1000) x 14 = 1.68 kWh
Energy = 19.2 kWh + 1.68 kWh = 20.88 kWh

## Slide 42: Basic AC Theory What is VA?
Power was measured in Watts. Power does useful work. The power that does useful work is referred to as
“Active Power.”
VA is measured in Volt-Amperes. It is the capacity required to deliver the Power. It is also referred to as the “Apparent Power.”
Power Factor = Active Power / Apparent Power
VA = E x I
PF = W/VA

## Slide 43: Basic Meter Math Power – VA
For a 120 Volt service drawing 13 Amps at 0.866 PF
How much power is being drawn?
Power = 120 x 13 x 0.866 = 1351 Watts
How many VA are being drawn?
VA = 120 x 13 = 1560 Volt-Amperes

## Slide 44: Basic Meter Math Power – VA
For a 480 Volt service drawing 156 Amps at 0.712 PF
How much power is being drawn?
Power = 480 x 156 x 0.712 = 53,315 Watts
How many VA are being drawn?
VA = 480 x 156 = 74,880 Volt-Amperes

## Slide 45: Basic Meter Math Power – VA
For a 120 Volt service drawing 60 Amps at 1.00 PF
How much power is being drawn?
Power = 120 x 60 x 1.00 = 7,200 Watts
How many VA are being drawn?
VA = 120 x 60 = 7,200 Volt Amperes

## Slide 46: Basic AC Theory Power Factor = 1.0

## Slide 47: Basic Energy Formula
The essential specification of a watthour meter’s measurement is given by the value Kh [ Watthours per disk revolution ]
The watthour meter formula is as follows:

## Slide 48
Meter Shop Accuracy Test Demo

## Slide 49: Questions to Answer
Why do we test?
How do we test?
What types of meter tests are there?
How do utility tests differ from customer request tests?
What is In-Service Testing?
How do we know meter tests are good?
What do we do with the test data?
What is the electrical theory behind the testing we do?

## Slide 50: Questions?
Tom Lawton
TESCO – The Eastern Specialty Company
Bristol, PA
215-785-2338
This presentation can also be found under Meter Conferences and Schools on the TESCO web site: www.tescometering.com