## Slide 1
ANSI C12 Standards Update
Prepared by Tom Lawton, TESCO
For Pennsylvania Rural Electric Association (PREA)
Tuesday, March 12, 2019
12:30 p.m. to 1:30 p.m.

## Slide 2: ANSI
American National Standard Institute, Inc.
  Not a government agency
  Standards do not have force of Law
  All compliance is voluntary
  ANSI doesn’t actually generate any standards
  Each standard is controlled by an industry organization as the “secretariat”
    For C12 NEMA (National Electrical Manufacturer’s Association) is the secretariat
    Paul Orr has been the NEMA’s secretary assigned to C12 for over ten years providing continuity to the process

## Slide 3: ANSI
American National Standard Institute, Inc.
  NEMA organizes committees to propose and review standards
  Standards are republished approximately every 5 years
  Standards codify consensus approaches in common practice
  Generally, they do not break new ground or deal in controversial issues
    This is changing. Can’t avoid issues any longer.

## Slide 4: ANSI C12
C12 Main Committee
  General makeup has expanded slightly over last few years
  34 voting members with representation from three groups:
    12 - Manufacturers: Meter, Socket, Test Equipment, etc.
    13 - Users: Utilities
    9 - General Interest: PUC, UL, IEEE, Consultants, etc.
  Meets twice a year in conjunction with EEI/AEIC Meter conference.

## Slide 5: ANSI C12
C12 Main Committee
  Has final approval for all activities on any C12 family standard
  Establishes Subcommittees (SC) and Working Groups (WG) to address various standards and issues
  Sub committees and Working Groups also meet twice a year in conjunction with EEI Transmission, Distribution and Metering Conference and also hold regular or ad hoc conference calls throughout the year as members put together drafts and other technical material for consideration at the next face to face meeting.

## Slide 6: ANSI C12
C12 Subcommittees
  Various subcommittees have been organized to review specific standards
  This is where the work is really done
  Each operates slightly differently
  Each meets on a schedule of its own choosing
  Most meet at EEI Biannual Transmission, Distribution and Metering Meetings
  Communication WG meets more often and longer
  Various subgroups meet frequently by teleconference

## Slide 7: ANSI C12 – Sub-Committees
Sub-Committee | Standards
C12 SC1 | C12.1, C12.4, C12.5, C12.10
C12 SC15 | C12.6, C12.7, C12.8, C12.9, C12.11
C12 SC16 | C12.20, C12.24
C12 SC17 | C12.18, C12.19, C12.21, C12.22, C12.23, C12.26
C12 SC29 | C12.29 Standard for Field Testing of Metering Sites
C12 SC31 | C12.31 VA Measurement Standard
C12 SC46 | C12.46 American National Standard for Electricity Meters - 0.1, 0.2 and 0.5 Accuracy Classes for the Measurement of Active, Apparent and Reactive Power

## Slide 8: ANSI C12 SC 1
C12.1– Code for Electricity Metering
  C12.1-2014 published in 2016
  Changes in this Revision
    Vector diagrams for balanced voltages are defined for common services
    Various definitions have been updated
    Bi-directional metering is added and tests updated to include bidirectional measurements
    Significant changes were made to the temperature rise section based on TRINIWOG tests

## Slide 9: ANSI C12 SC 1
C12.1 – Code for Electricity Metering
  Changes in this Revision (continued)
    Table 1b on maximum errors for reference standards was added
      Tightened to 0.05% and 0.02%
    Meter Accuracy Class 0.5 was added to the standard
    A new column was added to all tables with tighter specs for AC 0.5
    Section 5 on Standards for New and In Service Performance was completely rewritten
    Appendix D was rewritten and simplified
    A new Appendix F was added

## Slide 10: ANSI C12 SC 1
C12.1 – Code for Electricity Metering
  Safety
    Work is under way to move the safety related sections of C12.1 into C12.10
    C12 is coordinating efforts with UL and its work on Subject 2735
      UL2735C (Canadian) is about to be balloted
  New topics for next revision
    Changes to EMC Testing
    Addition of demand testing
    Revision of Temperature Rise Test based on additional testing

## Slide 11: ANSI C12 SC 1
C12.10 – Physical Aspects of Watthour Meters – Safety Standards
  Published C12.10 – 2011
  New revision in editorial review
    Will add safety sections from C12.1
    New safety requirements in development
    Working with UL to achieve consistency with UL 2735C
    10 new meter forms with auxiliary power pins

## Slide 12: ANSI C12 SC 1
New meter forms with Auxiliary Power

## Slide 13: ANSI C12 SC 1
New meter forms with Auxiliary Power

## Slide 14: ANSI C12 SC 1
New meter forms with Auxiliary Power

## Slide 15: ANSI C12 SC 1
New meter forms with Auxiliary Power

## Slide 16: ANSI C12 SC 15
C12.6 – Marking and Arrangement of Terminals for Phase-Shifting Devices used in Metering
  About to be reissued without change
C12.7 – Requirements for Watthour Meter Sockets
  Published – Feb 2015
Notes: 16

## Slide 17: ANSI C12 SC 15
C12.8 – Test Block and Cabinets for Installation of Self contained “A” Base Watthour Meters
  Published without change Dec 2012
  Will be deprecated once Figure 4 is moved to C12.7 in its next release
Notes: 17

## Slide 18: ANSI C12 SC 15
C12.9 – Test Switches and Plugs for Transformer Rated Meters
  Extensive Revision in last publication
  Specifications for test plugs added
  Published – March 2015
Notes: 18

## Slide 19: ANSI C12 SC 15
C12.11– Instrument Transformers for Revenue Metering, 10 kBIL through 350 kBIL
  Revised to parallel C57.13
  Published – July 2014
  New proposal on the table to revise barcode formats to allow QR and other more modern codes
Notes: 19

## Slide 20: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  C12.20 -2015 Published May 2017, Errata sheet also published
  0.1 Percent Accuracy Class Added
    New column added to all accuracy test tables
  Section 4.6 added to provide specifications for test outputs
    Type B port specifications were made “informative” due to a lack of unanimity. No Standard can force a manufacturer or utility to change their existing practices without their voluntary consent.

## Slide 21: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Tighter performance at reference conditions was specified
  Table 2 updated to make clear which forms and applications are Blondel compliant.
  Table 2A added to make clear which forms and applications are NOT Blondel compliant
    Therefore NOT covered by the standard

## Slide 24: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Polyphase loading for accuracy testing is allowed and details provided
    In next revision polyphase loading is proposed to be REQUIRED
  Non-Sinusoidal Performance Testing
    Harmonics WG
      Extensive performance testing using non-sinusoidal waveforms was added

## Slide 25: ANSI C12 SC16
C12.20 – 0.1, 0.2 and 0.5 Accuracy Class Metering
  Areas of activity for next revision
    Auxiliary device influences (requiring accuracy testing with any communication devices active)
    Disturbances test for memory corruption due to rapid power cycling such as during dropout/reclosure events
    Extended temperature operating ranges

## Slide 26: ANSI C12 SC29
Field Testing of Metering Installations
Field Testing of Metering Accuracy
  Hope to send for approval after next meeting
  Agreement on single set of acceptability criteria for testing under all three modes
    Voltage and current supplied by test equipment
    Site voltage, equipment provided current
    Site voltage and current (customer load)
  Final issue is review of “acceptability” values

## Slide 27: ANSI C12 SC31
VA and VAR Metering Standard
Establish a legal definition for VA and VAR
VA being addressed first in order to “Fast Track” the result.
  Definitions agreed upon
  Initial draft completed
At this meeting, new quantity “source VA” was proposed by Landis+Gyr
  Committee agreed to consider during a series of WEB conferences
Goal is to promote VA and active energy (watts) as the primary metering quantities.

## Slide 28: ANSI C12.46
Sub-Committee to Develop a Replacement for C12.1 and C12.20
  Adopts the structure of OIML IR-46
  Addresses 0.1, 0.2, and 0.5 Accuracy Classes
  Addresses Active, Reactive, and Apparent Energy
  Will be biggest change to metering standards in 100 years
  Review and editing of draft in process
  Expected Publication 2020

## Slide 29: Questions and Discussion
Tom Lawton
tom.lawton@tescometering.com
TESCO – The Eastern Specialty Company
Bristol, PA
215-785-2338
This presentation can also be found under Meter Conferences and Schools on the TESCO web site: www.tescometering.com